The Generic Patient Controlled Analgesia Pump Hazard Analysis

- System Domain

The system domain for the generic patient controlled analgesia (GPCA) pump is defined to include the following:

- Infusion pump

- Drug infusion set (bag / cassette / vial)

- Network

- Patient

- Environment (clinical setting, temperature, humidity, etc.)

- User / Medic

- Pump Components

- Infusion Module (Pump module)

  - Delivery mechanism

  - Bolus mechanism

  - Pump Sensors

  - Pump log

- User-Interface Module

  - Programming unit

  - Keypad

  - Pump console, audio unit

  - On-board Memory

- Error Handling Module

  - Alarm Handler

  - Power On Self Test (POST)

  - Watchdog timer

  - (Periodic) system checks

- Power Module

  - Power supply

  - A to D converter

  - Battery

  - Real Time Clock (RTC)

- Dose Error Reduction Module

  - Drug Library

- Communication Module (typically for PCUs or large volume external pumps)

  - Network controller

  - Channel communication

  - System (processor) bus

- Hazards

Hazardous or potentially harmful situations for the GPCA pump can be classified under the following categories

- Operational Hazards

- Environmental Hazards

- Electrical Hazards

- Hardware Hazards

- Software Hazards

- Mechanical Hazards (Physical Hazards)

- Biological and Chemical Hazards

- Use Hazards

A detailed hazard analysis is described in Section 4. Not all of these hazards are addressed by the (base) GPCA model. However, the model can be easily extended to incorporate additional functionality to address these hazards.

- Hazard Analysis

- Operational Hazards

| HID | Hazard | Pump Type | Cause | Action | Mitigated by | Safety Requirement |
| --- | --- | --- | --- | --- | --- | --- |
| 1.1 | Overinfusion | All | Programmed flow rate too high | Alarm(); Log() | Drug library | 1.1, 1.4.4, 1.4.11 |
| 1.2 | Overinfusion | All | Dose limit exceeded due to too many bolus requests | Alarm(); Log() | Flow sensor | 1.4, 3.4.6 |
| 1.3 | Overinfusion | All | (Programmed) Bolus volume/concentration too high | Alarm(); Log() | Drug library | 1.4, 3.4.6 |
| 1.4 | Overinfusion/ Underinfusion | All | Incorrect drug concentration specified | Alarm(); Log() | Barcode scanner | 1.1, 6.1.3, 6.1.4 |
| 1.5 | Underinfusion | All | Programmed flow rate too low | Alarm(); Log() | Drug library | 1.1, 6.1.3, 6.1.4 |
| 1.6 | Underinfusion | FRN | Air in line | Alarm(); Log() | Flow sensor | 1.9 |
| 1.7 | Underinfusion | FRN | Occlusion (supply side and patient side) | Alarm(); Log() | Flow sensor | 1.10 |
| 1.8 | Underinfusion | FRN | Reservoir empty | Alarm(); Log() | Flow sensor; Drug library | 1.5 |
| 1.9 | Underinfusion | FRN | Reservoir low | Alert(); Log() | Flow sensor; Drug library | 1.5 |
| 1.10 | Underinfusion | All | Flow rate does not match programmed rate | Alarm(); Log() | Flow sensor | 1.2, 6.1.3, 6.1.4 |
| 1.11 | Deflation issue | FRN | Inability of device and/or device components to release gas or air | Alert(); Log() |  |  |
| 1.12 | Filling problem | All | Inability to Auto fill | Alert(); Log() |  |  |
| 1.13 | Improper flow | FRN | Free flow of drug | Alarm(); Log() | Flow sensor | 1.2.2 |
| 1.14 | Improper flow | FRN | Bleed back; Reflux within device | Alarm(); Log() | Flow sensor | 1.8 |
| 1.15 | Improper flow | FRN | Fluctuation of Tidal Volume | Alarm(); Log() |  |  |
| 1.16 | Improper flow | All | Inaccurate flow rate; Infusion intermittent | Alarm(); Log() | Flow sensor | 1.2 |
| 1.17 | Inflation issue | FRN | Inability of device and/or device components to expand or enlarge with gas or air | Alert(); Log() |  |  |
| 1.18 | Low Pressure | All | Decrease in Pressure; No Pressure | Alarm(); Log() |  | 1.10.3, 1.10.4, 1.10.5 |
| 1.19 | High Pressure | All | Increase in Pressure | Alarm(); Log() |  | 1.10.3, 1.10.4, 1.10.5 |
| 1.20 | Low Pump speed | All | Decreased pump speed; Pumping stopped | Alarm(); Log() | Flow sensor | 1.1.5, 1.1.8, 1.2.3 |
| 1.21 | High Pump speed | All | Increased pump speed | Alarm(); Log() | Flow sensor | 1.2.3 |
| 1.22 | Failure to alarm | All | Defective alarm unit; Delayed alarm detection | Log() |  |  |
| 1.23 | False alarm | All |  | Log() |  |  |
| 1.24 | Failure to prime | FRN | Air in line | Alert(); Log() | Flow sensor | 1.9 |
| 1.25 | Incorrect therapy | FRN | Prescription/dosage values fall out of default value range | Alert(); Log() | Drug library; Barcode scanner | 5.1 |
| 1.26 | False alarm | FRN | Inappropriate prompts | Log() |  |  |
| 1.27 | Air bubble introduced in blood stream | All | Air in line | Alarm(); Log() | Flow sensor | 1.9 |

- Environmental Hazards

| HID | Hazard | Pump Type | Cause | Action | Mitigated by | Safety Requirement |
| --- | --- | --- | --- | --- | --- | --- |
| 2.1 | Failure to operate/ Pump malfunction | All | Temperature /Humidity/ Air pressure too high or too low |  |  | 7.1 |
| 2.2 | Contamination | FRN | Contamination due to spillage / exposure to toxins |  |  |  |
| 2.3 | Incorrect therapy | FRN | Patient is underweight; Patient is overweight; Patient has medical condition that disallows use of specific pump | Alert(); Log() | Barcode scanner | 5.1.1 |
| 2.4 | Failure to attend alarm | All | Background noise (may cause alarms not being heard by medic) |  |  | 3.2.3 |
| 2.5 | Failure to attend alarm | FRN | Patient muffles alarm (ambulatory/portable pump) |  |  | 3.2.3 |
| 2.6 | Failure to attend alarm | FRN | Inaudible or no voice prompts |  |  | 3.2.3 |
| 2.7 | Tampering | FRN | Patient tampers with pump settings without authorization |  |  | 2.1 |
| 2.8 | Tampering | FRN | Panel lock broken or opened during infusion | Alert(); Stop() |  | 2.1, 3.3 |
| 2.9 | Tampering | FRN | Panel/door opened during infusion; Infusion started when door open | Alert(); Log() |  | 2.1, 3.3 |
| 2.10 | Interference | All | Electrical interference from cell phones, ESD etc. |  |  | 6.1 |
| 2.11 | Interference | FRN | Inadequate shielding provided |  |  | 6.1 |
| 2.12 | Overheating | FRN | Fire |  |  | 7.1.2 |
| 2.13 | Contamination | FRN | Battery leak |  |  |  |

- Electrical Hazards

| HID | Hazard | Pump Type | Cause | Action | Mitigated by | Safety Requirement |
| --- | --- | --- | --- | --- | --- | --- |
| 3.1 | Overheating | FRN | Incorrect or loose interconnections between devices – channel error; | Alarm(); Log() |  | 7.1.2 |
| 3.2 | Overheating | FRN | Supply processor charge too high; Insufficient cooling/faulty heat sink; Unintended magnet quench | Alarm(); Log() |  | 7.1.2, 7.3 |
| 3.3 | Charge Error | All | Battery could not be charged | Alarm(); Log() |  | 4.1.8 |
| 3.4 | Supply Voltage Error | FRN | Supply voltage too high; Supply voltage too low; Battery voltage exceeds limits |  |  | 7.3 |
| 3.5 | Battery Failure | FRN | Battery voltage too low; Battery depleted | Alarm(); Log() |  | 4.1 |
| 3.6 | A-to-D conversion Failure | All | A-to-D conversion failed |  |  |  |
| 3.7 | Electric shock | FRN | Leakage Current too high (pump could be source of electric shock) |  |  | 4.2.1 |
| 3.8 | Electric shock | FRN | Electrical power failure; Power surge |  |  | 4.1.9 |
| 3.9 | Electric shock | All | Inadequate resistance; Loss of resistance |  |  |  |
| 3.10 | Circuit failure | FRN | Electrical shorting; High impedance; Low impedance |  |  | 4.1.9 |
| 3.11 | Electromagnetic compatibility issue | FRN | Electromagnetic interference; Electrostatic discharge; Radiofrequency interference |  |  |  |

- Hardware Hazards

| HID | Hazard | Pump Type | Cause | Action | Mitigated by | Safety Requirement |
| --- | --- | --- | --- | --- | --- | --- |
| 4.1 | System failure | All | Malfunctioning component | Alarm(); Log() |  | 3.3, 3.5 |
| 4.2 | System failure | FRN | System malfunction – RTC not synchronized (date/time register not same as the RTC); Clock frequency check failed |  |  | 3.3, 3.4 |
| 4.3 | System failure | All | CPU test failed; One or more of the system component failed |  |  | 3.3, 3.4 |
| 4.4 | System failure | All | Synchronization error between pump components |  | Drug library | 3.3.4 |
| 4.5 | Channel error | FRN | synchronization problem between channels on a multi-channel pump |  |  | 3.5 |
| 4.6 | Network error | FRN | Pump not compatible with networked / integrated device |  |  | 3.3.4 |
| 4.7 | Memory failure | FRN | System malfunction – RAM test failed; Attempted write to memory failed; Critical value data integrity error |  |  | 3.3, 3.4 |
| 4.8 | Memory failure | FRN | System malfunction – ROM (or external flash memory) CRC test failed |  |  | 3.3, 3.4 |
| 4.9 | Watchdog failure | All | System malfunction – Watchdog timer test failed; Watchdog not interrupted in 90 seconds |  |  | 3.4.4.5 |
| 4.10 | False alarm | All | False watchdog interrupt |  |  |  |
| 4.11 | Incorrect test results | All | False negative test result; False positive test result; Incorrect measurement; Test result inaccurate |  |  |  |
| 4.12 | Incorrect dose value entered | FRN | Key debounce not detected |  | Drug library | 2.3 |
| 4.13 | Failure to alarm | All | Sensor failure |  |  |  |

- Software Hazards

| HID | Hazard | Pump Type | Cause | Action | Mitigated by | Safety Requirement |
| --- | --- | --- | --- | --- | --- | --- |
| 5.1 | Data error | FRN | Failure to backup; Data retrieval error Could not write to pump log |  |  | 1.7.1 |
| 5.2 | Data error | FRN | Unable to retrieve data from drug library; Failure to transmit record |  |  |  |
| 5.3 | Incorrect version | FRN | Software updates not installed; Incorrect version installed |  | Barcode scanner | 5.1.8 |
| 5.4 | Failure to alarm | All | Communication problem between channels | Log() |  |  |
| 5.5 | Pump could not be silenced | FRN | Alarm priority set incorrectly | Log() |  |  |
| 5.6 | Incorrect dose administered | FRN | Incorrect drug library; Old version of drug library |  |  | 5.1.8 |
| 5.7 | Channel error | FRN | Failure to recognize new channels added to pump |  |  |  |
| 5.8 | Communication error | All | System malfunction – RF Communication test failed |  |  | 3.3.4 |
| 5.9 | Pump failed to startup | FRN | One or more of the POST tests failed |  |  | 3.4.5 |
| 5.10 | Pump failed to shut down | All | Failure to auto-stop (following a critical failure that requires pump to be stopped) |  |  |  |
| 5.11 | Pump reverts to default dose values | All | Programmed dose set incorrectly; Inappropriate reset to default |  | Drug library | 5.1.3 |
| 5.12 | Incorrect test results | All | False negative test result; False positive test result; Incorrect measurement; Test result inaccurate |  |  |  |

- Mechanical Hazards (Physical Hazards)

| HID | Hazard | Pump Type | Cause | Action | Mitigated by | Safety Requirement |
| --- | --- | --- | --- | --- | --- | --- |
| 6.1 | Unable to set dose, start/ stop/ reset pump, silence alarm | FRN | Broken part (e.g., broken keypad) | Alert() |  | 3.3 |
| 6.2 | Incorrect dose value entered | FRN | Key stuck / depressed | Alarm() |  | 2.3 |
| 6.3 | No alarm signal | FRN | Speaker / Audio unit failure | Log() |  | 3.3 |
| 6.4 | Physical Damage to pump | All | Falling; Shear; Stress |  |  |  |
| 6.5 | Injury to medic/patient | FRN | Sharp edges |  |  |  |
| 6.6 | Pump stops infusion | All | Pump motor fails; Pump unable to stroke |  | Flow sensor | 3.5 |

- Biological and Chemical Hazards

| HID | Hazard | Pump Type | Cause | Action | Mitigated by | Safety Requirement |
| --- | --- | --- | --- | --- | --- | --- |
| 7.1 | Biological/Chemical Hazard | FRN | Device contaminated during use; Device contaminated by blood/leaking fluid |  |  |  |
| 7.2 | Biological/Chemical Hazard | FRN | Inadequate device cleaning; Residue after contamination; Failure to flush; Failure to disinfect |  |  |  |

- Use Hazards

| HID | Hazard | Pump Type | Cause | Action | Mitigated by | Safety Requirement |
| --- | --- | --- | --- | --- | --- | --- |
| 8.1 | Overfill | All | Incorrect fill volume specified | Alert(); Log() | Barcode scanner |  |
| 8.2 | Short fill | All | Incorrect fill volume specified | Alert(); Log() | Barcode scanner |  |
| 8.3 | Knowledge-based failure | All | Operating instructions incomplete; Inaccurate labeling |  | Barcode scanner |  |
| 8.4 | Knowledge-based failure | FRN | Medic fails to recognize hazardous situation |  |  |  |
| 8.5 | Knowledge-based failure | FRN | Pump does not display adequate dosage information on the display |  |  |  |
| 8.6 | Rule-based failure | FRN | Incorrect prescription given to patient; Incorrect drug library loaded |  | Barcode scanner | 2.2, 5.1.8 |
| 8.7 | Overinfusion | FRN | User / Patient change infusion settings inadvertently |  | Drug library | 2.2, 5.1 |
| 8.8 | Underinfusion | FRN | User / Patient change infusion settings inadvertently |  | Drug library | 2.2, 5.1 |
| 8.9 | Patient incapacitated | FRN | Home care patient unable to administer dosage/service alarm condition |  |  |  |
| 8.10 | Attentional failure | All | Incorrect prescription entered |  |  | 2.2, 5.1 |
| 8.11 | Memory failure | FRN | Too few doses administered; Patient given multiple doses |  | Flow sensor |  |
| 8.12 | Incorrect dose settings | FRN | Key pressed too long | Alarm() |  | 2.3.1 |
| 8.13 | Inadequate training | All | User not trained to use pump; User not familiar with pump |  |  |  |
| 8.14 | Incorrect dose mode | FRN | Incorrect units used for specifying dose parameters (e.g., ml/hr instead of mcg/hr) |  | Barcode scanner |  |