Solving the Problem of Duplicate Medical Device Unique Identifiers

1. Introduction

Unique identifiers are generated by medical device systems to individually label different instances of clinical information objects, for example different patients, different studies, different reports, and different images.  In order to protect patient safety, it is absolutely essential that no two different entities ever be labeled with the same unique identifier.  It has been the federal government’s experience, however, that occasionally the same unique identifiers are generated for different medical information objects.  This is caused by poorly designed unique identifier generation methods that fail due to computer malfunction and/or human error.  This is a very serious problem that can affect patient safety.

2. Background

Digital Imaging and Communications in Medicine (DICOM) was the first medical standard to use unique identifiers.  DICOM uses unique identifiers (UIDs) throughout to identify different instances of clinical information objects. When a duplicate UID is generated, different studies or different images are assigned the same UID value, and it is possible to associate the study or image with the wrong patient.  It is estimated that the frequency of occurrence of duplicate UIDs in DICOM is on the order of magnitude of one in a thousand.

HL7 Version 3.0 uses an analogous unique identifier called an Object ID (OID) to label its information object instances.  As similar factors that are present in DICOM implementations are also present in HL7 ones, it is expected that OIDs will also have a duplication problem.

3. Root Cause of Problem

The root cause of the DICOM/HL7 duplicate unique identifier problem is that while the standards define the format and rules for the unique identifiers (that is, they are both ISO 8824 Object Identifiers), the vendors are free to choose their method of implementing them.  Many vendors devise private methods that fail in practice.  These cause problems with some current implementations.

3.1 Problems Using Natural Primary Keys to create the Unique Identifier

Natural primary keys have business meaning because they contain attributes or a combination of attributes that are unique by virtue of business semantics.  For example, an organization’s HL7 OID may be designed to contain a facility id, a patient id, and a study id.  Natural keys are almost always a problem because organizations inevitably change their business rules. For instance, an organization might merge with another and have new business rules requiring it to use different identifier schemes.  As a result, the new OIDs that are created might be identical to those generated by the old method or to those that are of another facility.

3.2 Problems Using a Synthetic value to create the Unique Identifier

In order to create unique identifiers that have no business meaning, a synthetic value should be used.  While there are some very good methods for creating synthetic values in the computer industry, a number of unsatisfactory methods are common in medical systems.

3.2.1 One-Up Counter

The simplest synthetic value is based upon the one-up counter.  The application uses a monotonically increasing number to generate unique identifiers like xxx.1, xxx.2, xxx.3, etc.  The counter is maintained in a disk file on the device.  Whenever there is a disk glitch, the counter can be reset.  It is reset whenever the disk is replaced.  The counter may be reset when a new version of the software is loaded.  In some cases, a vendor’s technician resets the counter as part of normal preventive maintenance.

3.2.2 Resetting Local Computer Time

The unique identifier can incorporate the local date and time values.  Duplicate unique identifiers can be generated if the clock is set back.  This may happen manually when an operator observes an incorrect time setting and resets the clock.  It may also occur automatically at the termination of daylight savings time each fall when clocks are turned back one hour to return to standard time.  Duplicate UID problems can also arise when the system clock malfunctions due to a hardware/software problem.

3.2.3 Same Identifiers at Multiple Sites

Vendor generated unique identifiers in the same namespace at different local sites.  As a result, duplicate values exist across the enterprise.  This problem will be a cross-enterprise interoperability barrier for organizations such as the VA/DoD that have data consolidation and sharing requirements.  You don’t want duplicates anywhere really, not just in what you define as the “enterprise”.

3.3 Resultant Problem – Dealing with existing Duplicate UIDs

There are over a billion DICOM images archived world-wide. Maybe as many as a million of them have duplicate UIDs.  How should end-user systems deal with instances of duplicate unique identifiers?   This currently is a big problem.  When a duplicate UID occurs, it must be detected and properly handled by the Picture Archiving and Communications System (PACS).  However, PACS vary in their handling of objects with duplicate UIDs.  Some PACS are designed with the assumption that all UIDs are unique.  These PACS could reject images with duplicate UIDs, although this may cause a limited disruption of service and a possible loss of data.  Other PACS might not detect the error and associate the images with the wrong patient and study.  Another PACS might mask the problem entirely and use other information to associate the images with the correct patient and study.  The problem may then surface later when these images are passed to another DICOM system, one that might operate in a different way and miss-associate the images with the wrong patient or study.

4.0 Challenges in Achieving a Solution

There are a number of diverse federal/public/commercial stakeholders that should be organized and involved to solve this critical medical device issue.  It will be a coordination challenge to utilize and engage the federal government (for example NSF, NSA, NIST, VA/DOD, FDA, HHS), standards associations (i.e. NEMA, ISO, HIMSS, RSNA, DICOM, HL7, and The Open Group), medical device vendors (i.e. PACS, RIS, HIS, speciality systems, modalities) and academia to address this issue.  The HCMDSS workshop and NITRD will be important in facilitating and guiding this discussion, and resolving issues such as Data Stewardship.

5.0 Important research needs

Finding a ‘best solution’ to solve duplicate unique identifiers requires thorough research by a committee of diverse stakeholders working through the risks and impacts of implementation.  NIST should provide technical guidance.

There currently exist multiple models for universally unique id generation.  In particular, the use of natural vs. synthetic primary keys should be explored. Also the use of a decentralized vs. centralized model could be analyzed. Research into performance limits may be useful in guiding implementation. Any identifier schemes that pertain to patient safety should have fail-safe structures built into them. These structures might include check digits and random components to assure the uniqueness and integrity of the identifier.

There may be implementation issues imposed by limits of the current DICOM and HL7 standards (64 characters numeric ISO 8824 Object Identifiers).  The committee of stakeholders will need to determine the best way to generate the unique identifiers.  This may impact the DICOM and HL7 standards.

The flip side of the problem is to determine how systems should deal with the duplicate unique identifiers in new data and in the legacy data.  What is the best design for a system that has to handle duplicate unique identifiers?  How should a system handle objects with duplicate unique identifiers when it encounters them?

Once the ‘best’ solutions have been determined, the committee will need to ensure that organizations are willing to accept the costs of upgrading their software.  This requires sufficient regulatory compliance, so the FDA should be involved.  Marketplace persuasion similar to that used by HIMSS and RSNA for IHE Compliance will play an important role.  The certification of medical devices generating unique identifiers and handling duplicates will be critical to solving the issue of non-compliance.

6.0 Roadmap for the next 5 to 10 years

There are a number of milestones on the roadmap that will guide the solution.  The first milestone will be educating the diverse stakeholders on the criticality of the issue. The issue of data stewardship will be addressed.

The second milestone will be the selection of effective methods for generating universally unique identifiers and the standardization of approved algorithms and approach.  This standards document may include choosing the “best” from multiple algorithms (such as time-based, name-based, and randomly generated). The approved solution will need to include metrics on how we can verify the solution.  The second milestone will also include the determination of “how best” to deal with the problem of the duplicate unique identifiers in the data.

The third milestone will be the demonstration of the validity of the standard and the verification of test tools.

The fourth milestone will be the VA/DoD and other customers requesting and procuring a certified solution. Vendors will develop to the requirements and certify their solution with the committee approved guidelines.

The fifth milestone will be the deployment of the solution and resolving implementation issues.  This includes determining migration strategies and the data cleansing process.

Biography

Peter Kuzmak is a Computer Specialist with the VA’s VistA Imaging Project.  He is responsible for DICOM issues.

Andrew Casertano provides technical support to the DOD’s Military Health System on DICOM, HL7 and IHE.

Daniel Carozza, MD consults to the VA on regulatory compliance and DICOM issues.

Ruth Dayhoff, MD is National Director of the VA’s VistA Imaging Project.

Keith Campbell consults to the VA on electronic health record issues.